Safe use of electrosurgery in gynaecological
laparoscopic surgery
- Applied physics
- Mechanism of electrosurgery
  - electromagnetic energy  o thermal energy
  - Various tissue effects of electrosurgery
Temperature (°C)	Tissue effect	Use
37		Normal temperature
50	Cell death	Irresversible damage
60–80	Denaturation of proteins and collagen / Vaprisation of fluid and shrinkage of coagulum	Coagulation / Desiccation
100	Vaporisation and explosion of cells	Cutting
200	Tissue charring	Carbonisation

- Fulguration
  - is non-contact + high voltage sparking
  - Superﬁcial layer of black Coagulation
  - The high voltage
    - helps to overcome impedance of the intervening air between the active electrode and the tissue.
    - This increases the risk of stray current burns.
- Monopolar electrosurgical instruments
- The ART of monopolar instrumentation
  - works by concentrating the current (increasing its density) at the active electrode to produce the desired thermal tissue effect and dispersing it at the dispersive electrode to prevent unintended tissue burns.

- Factors modifying electrosurgical tissue Effects
  - Waveform
  - Tissue contact
	cut	‘coagulation’
Contact	White coagulation / Coagulation	White coagulation / Coagulation
Non-contact	Clean cut / Cut	Superficial black coagulation / Fulgration
	Continuous low-volt	Interrupted high-volt
	rapid rise in tissue temperature to more than 100°C, which vaporises or cuts tissue with minimal coagulation	a slow increase in temperature to less than 100°C, which desiccates and coagulates tissue.

  - Power output
    - In Watts
    - Recommended:
      - lowest effective power setting
      - cut mode
        - 50 W and 80 W
      - coagulation mode
        - 30 W and 50 W.
    - muscular patients require lower settings compared with obese or emaciated patients.
  - Electrode surface area
    - The smaller the electrode, the higher the current concentration.
    - Should exploit this variable to achieve the desired tissue effects without increasing the power setting
  - Activation time
    - Long activation time increases the extent of tissue damage, whereas too short a time may result in inadequate tissue effect.

  - Tissue impedance
    - Low impedance tissues:
      - muscles and skin
    - High impedance tissues:
      - scarred tissue and fat
    - Eschar has high impedance to current
      - Cleaning the active electrode of eschar reduces impedance and enhances the electrosurgical effect.
- Conventional bipolar devices
  - two electrodes situated at the tip of the instrument
  - generally safer than monopolar instruments
  - vessel-sealing effect through
    - mechanical compression +
    - electrosurgical effects of desiccation and coagulation
  - Mushroom (outside loop) effect
    - Mechanism:
      - Grasped tissue desiccates and coagulates 
      - Impedance increases 
      - forcing the current to take a path of least impedance outside the jaws of the bipolar instrument.
  - Electrical bypass effect
    - Mechanism:
      - Overcompression of the tissue between the jaws of the bipolar instrument may cause them to touch 
      - leading to electrical bypass 
      - deﬁcient tissue coagulation
- Advanced bipolar devices
  - See tog article (advanced energy sources)

- Complications
  - Introduction:
    - Incidence of laparoscopic electrosurgical injuries:
      - 2–5 per thousand procedures
    - Incidence of unrecognized injuries:
      - 70%
    - The delayed diagnosis of these complications is associated with increased morbidity and mortality
  - Direct application
    - Definition:
      - Result from lateral thermal spread to unintended tissue near the tip of the electrosurgical instrument during activation
    - Extent of lateral thermal spread increase with:
      - Current:
        - Continuous
      - Voltage
        - High
      - Power setting
        - Higher setting
      - Tissue compression
        - Low
      - Pedicle size:
        - Big pedicle
      - Application time
        - Longer application
      - Instrument type
        - Monopolar coagulation
    - Monopolar devices > bipolar > ultrasonic ones
    - Pedicle effect
      - monopolar instrument is applied to a structure with a narrow vascular pedicle or adhesion.
      - unintended burn occurs at the remote narrow pedicle, or adhesion where the current density is higher.
    - Prevention:
      - To secure hemostasis near such structures, use sutures, clips or staples rather than electrosurgical devices.

  - Inadvertent activation
    - Prevention:
      - Avoid accidentally stepping on the foot pedal.
      - When the active electrode is not in use:
        - remove it from the body and
        - place it in a dry, rigid plastic holder
      - Use an audible activation tone to be heard by the team
  - Residual heat
    - Energy devices can maintain the heat at their tips for a variable time after deactivation.
    - ultrasonic energy instruments have higher residual heat than electrosurgical instruments
    - Prevention:
      - Avoid touching vital structures with the tip of an electrosurgical device immediately after deactivation.
      - If the bowel is inadvertently touched with a hot device, it should be examined for blanching and suturing should be considered to avoid delayed perforation.
  - Insulation failure
    - breakdown of the insulation layer around the active electrode.
    - incidence
      - 20% in reusable instruments
      - 3% in disposable instruments
    - Distal third of the instrument being the most commonly affected site
    - counterparts.
    - Causes of breakdown:
      - Repeated cleaning and sterilization
      - normal wear and tear
      - Use of high-voltage output
    - Prevention:
      - Inspect the instrument before use
        - Unfortunately most of these defects are not visible to the naked eye.
      - Active electrode with an indicator shaft was designed with two layers of insulation (black outer and yellow inner)
      - Active electrode monitoring (AEM) technology prevents stray current burns from insulation failure and capacitive coupling.

  - Antenna coupling
    - active electrode cord (transmitting antenna) emits electromagnetic energy in the air, which is captured by a nearby inactive cord or wire (receiving antenna).
    - A type of capacitive Coupling
    - Prevention:
      - Separating the laparoscopy tower from ESU
      - Avoiding parallel arrangement of cords
      - Lowering the power setting
  - Direct coupling
    - When the active electrode touches another metal instrument such as a suction irrigator or camera telescope.
    - Current from the active electrode ﬂows to the second instrument and potentially burns any tissue it touches.
    - Prevention:
      - energy must not be activated until the instrument is out of the metal trocar and its tip is in view.
      - Ports must be placed so as to avoid instrument shafts from touching the bowels.
      - The active electrode and other metal instruments should be kept in panoramic view
      - Surgeon should be the only person to activate the energy.
      - In the event that an arc to an adjacent instrument is seen
        - The surgeon should examine the length of that instrument, looking for any contact with vital tissue.
        - If there is evidence of a burn in that tissue, corrective measures such as suturing should be taken
  - Capacitive coupling
    - transfer of electric current from the active electrode, through intact insulation, into adjacent conductive materials without direct contact
    - Examples of situations in which capacitive coupling can pose a risk to the patient include:
      - When the active electrode is passed down a metal suction irrigator, an operative laparoscope or a metal cannula with a plastic gripper (hybrid cannula).
      - When the insulated shaft of the active electrode touches non-targeted tissue such as bowel or adhesion.
      - When the active electrode induces a current in a nearby cold instrument.
      - When current is induced into adjacent tissue and instruments in single-port laparoscopy.
    - Prevention:
      - Avoiding hybrid cannulas
      - Lowering esu power setting
      - Using the ‘cut’ rather than ‘coagulation’ mode
      - Using short interrupted activation
      - Avoiding open activation
      - Operating close to metals in the operative ﬁeld.
      - Adaptive electrosurgical technology
        - reduces capacitive coupling but has no effect on insulation failure.
      - AEM technology
        - eliminates injuries caused by both insulation failure and capacitive coupling
- Technological developments and safe Electrosurgery
  - Unintended alternate site burns:
    - ground point burns
    - dispersive electrode burns
  - Prevention of unintended alternate site burns
    - Isolated ESUs
    - Contact quality monitoring (CQM)
  - Active electrode monitoring
    - Conventional monopolar device
      - Outer insulation layer covers the shaft of the active electrode.
    - AEM instruments
      - two extra coaxial layers:
        - conductive (protective) shield
        - second outer insulation
    - AEM system
      - continuously monitors the conductive shield for stray currents caused by insulation failure and capacitive coupling..
      - protective shield is considered as a second dispersive electrode, which returns stray currents safely to the AEM monitor and then back to the ESU.
      - If the AEM monitor detects a dangerous level of stray energy (about 2 W), it deactivates the ESU to prevent tissue burns.
- Electrosurgery in single-port laparoscopy
  - single-port laparoscopy (SPL), in which three or four instruments are passed through one port
  - heightened awareness of the potential risks of monopolar instrumentation (insulation failure, direct coupling, and capacitive coupling).
  - Attributed to
    - Increased length of zone 2
    - Proximity and crossing of instruments (‘sword ﬁghting’)
  - To reduce stray current injury in SPL:
    - use alternative devices such as bipolar or ultrasonic instruments.
    - with monopolar instruments:
      - Use those with AEM technology.
      - Use a metal cannula to disperse capacitive charge into the abdominal wall.
      - Use a 3-mm device to reduce capacitive coupling.
- Electrosurgery and electromagnetic Interference
  - Electrosurgery can interfere with cardiac implantable electronic devices (CIED) such as:
    - permanent pacemakers (PPM)
    - implantable cardioverter deﬁbrillators (ICD),
  - Consequences:
    - damage or inhibit the CIED device
    - burn the myocardium
    - Arrhythmias and asystole
  - Prevention:
    - Use bipolar or ultrasonic devices in patients who are highly dependent on CIED.
    - When using monopolar instruments:
      - Place the dispersive electrode as far away as possible from the CIED.
      - Avoid the capacitively coupled return electrode.
      - Use lower power, cut mode to coagulate and short activation.
      - Avoid current vector crossing the CIED.
      - Deactivate ICDs just before surgery, then activate after surgery
      - Monitor PPM with ECG during surgery and reprogram after surgery if needed.

- Electrosurgical smoke
  - Disadvantages:
    - Reduces laparoscopic visualisation,
    - potentially carcinogenic chemicals
    - viruses.
    - irritation of the eyes and upper respiratory tract
  - Prevention:
    - Smoke evacuation systems should be used to reduce the above risks.
    - Surgical masks are ineffective because they only ﬁlter particles down to 5 lm, whereas 77% of smoke particles are smaller than or equal to 1.1 lm.

- Good practice in electrosurgery
  - Monopolar devices
    - Preven skin burns:
      - Return electrode monitoring
      - Make sure dispersive electrode is well applied
      - Do not apply the dispersive electrode over
        - Bony prominences
        - Metal prosthesis
        - Scar tissue
        - Hairy skin
        - Pressure areas.
    - Prevent capacitive coupling
      - Avoid open activation.
      - Do not use in tissues with metal clips or staples in situ
      - Do not use hybrid cannulas
    - Prevent direct coupling
      - Avoid activation in close proximity to or in contact with another metal instrument.
    - Avoid insulation failure and skin burn
        - Active electrode monitoring.
        - Examine outer sheath of monopolar instruments for break-ups before every operation.
  - For both bipolar and monopolar devices
    - Minimize thermal spread
      - Use short intermittent activations.
      - Use the continuous low-voltage waveform ‘cut’ mode for contact coagulation.
      - Use the lowest possible power setting.
      - Allow a safety margin.
      - Terminate at the vapor phase.
      - Irrigation with fluid before reactivation.
      - Avoid over-compression of grasped tissue to prevent the bypass effect
      - Vary the surface area of the active electrode to achieve the desired effect without increasing the power setting.
    - For good coagulation:
      - Avoid tension on the tissue during activation
      - Keep the jaws of the instrument clean
      - do not include a big bundle of tissue in the jaws
      - Consider skeletonization vessels before application
    - Tissue stuck:
      - Avoid:
        - Release the tissue just before current ﬂow
      - Manage:
        - When stuck to tissue, re-approximate the jaws and reactivate before opening them.
    - In patients with comorbidities such as liver cirrhosis, prolonged steroid use, atherosclerosis, diabetes, malnutrition and collagen diseases, be extra cautious and consider alternative surgical methods because these conditions may affect the blood vessels